LPRA RAIL SCALE SPECIFICATIONS

170 Ton Capacity - (1) 13ft Deck, (1) 26ft Deck, Hydraulic, Modular

GENERAL PROVISIONS:

Furnish and install a modular steel platform railroad track scale and associated electronic controls.

The scale shall be a Cardinal Scale model H1326-80LPRA-100.

The scale shall have a profile not to exceed 36 inches, which is measured from the top of the scale to the top of the foundation slab or pier at the load cell bearing points.

The scale shall consist of one 13ft module and one 26ft module with a dead space between modules.

The scale shall be designed to function as either a single standalone platform or a multi-platform system.

The scale shall have a gross weighing capacity of 170 tons and shall have a section capacity of 85 tons.

The scale shall be designed to accommodate 115lb rails and have an E-80 Cooper AREMA Rating.

The load cells shall be constructed of stainless steel.

The scale shall be NTEP certiﬁed and shall meet the requirements as set forth by the National Institute of Standards and Technology (NTEP) Handbook 44 current edition for Class IIIL devices. The scale manufacturer shall provide an NTEP Certiﬁcate of Conformance attesting to conformance with these standards.  Provisional certiﬁcation will not be accepted.

The scale shall be Measurement Canada (MC) certified and shall carry the appropriate MC Approval. The scale manufacturer shall provide a Measurement Canada Approval number attesting to conformance with these standards.  Provisional certiﬁcation will not be accepted.

The scale shall be designed to comply with the requirements of the AAR Rail Scale Handbook.

The design and manufacture of the scale weighbridge, load cells, digital instrumentation and associated accessories shall be of one manufacturer to maximize compatibility and availability of components.

The manufacturer shall provide with the bid proposal a listing of major spare parts and their prices including, but not limited to, replacement load cells, weight indicator, circuit boards and associated accessory parts.

SCALE FOUNDATION REQUIREMENTS:

The foundation shall meet all local requirements and the minimum speciﬁcations as stated herein.

The minimum soil bearing capacity shall be 4,000 psf. The buyer shall be responsible for determining whether or not the soil condition is adequate.

The foundation shall extend 12” past the frost line at the load cell piers.

The foundation shall be constructed of concrete with a minimum strength of 3,000 psi at a 28-day cure.

The foundation shall be reinforced in all load-bearing areas. The reinforcing steel shall be 60 KSI yield strength and conform to ASTM A615 grade 60 minimum.

The foundation shall be constructed such that positive drainage away from the foundation is maintained.

The foundation shall be designed to include two approaches, one at each end of the scale, in accordance with local regulations and guidelines of the National Institute of Standards and Technology Handbook 44, current edition.

WEIGHBRIDGE SPECIFICATIONS:

The weighbridge shall consist of prefabricated steel deck modules.

There shall be no field welding required for the installation of the scale.

Each weighbridge module shall be constructed of W18 I-beams spanning the full length.

SURFACE PREPARATION AND FINISH:

The structural steel shall be shot blasted to an SSPC-SP10 condition to remove rust and mill scale, then protected with a baked-on, UV-protected, heavy-duty polyester anticorrosion industrial tan powder paint for the highest quality and durable ﬁnish.

LOAD CELL SPECIFICATIONS:

All load cells shall be of hydraulic compression design and shall have a minimum capacity of 100,000 pounds with a 150% of capacity overload rating.

Load cells shall be certiﬁed by NTEP and shall meet the speciﬁcations as set forth by the National Institute of Standards and Technology Handbook 44 for Class IIIL, multiple cells, 10,000 divisions. The manufacturer shall provide a NTEP Certiﬁcate of Conformance attesting to compliance with these requirements.

The load cell shall be constructed from stainless steel and shall be environmentally sealed to an IP68 / NEMA 6P rating.

Load cells shall output a hydraulic pressure signal to a corresponding pressure transducer at the totalizer junction box.

Temperature compensation shall take place within the pressure transducer circuitry.

The load cell shall be provided with sufficient tubing to reach the totalizer junction box.

The load cell shall be manufactured by the scale manufacturer and shall be a Cardinal Scale model SST100.

TOTALIZER JUNCTION BOX AND TUBING:

The outputs from the pressure transducers shall be electrically summed with provisions for adjustment of individual load cells provided. Totalizers employing the mechanical summation of load cell pressures and / or utilizing shims are not acceptable.

The totalizer junction box shall be mounted adjacent to the location of the weight-indicating instrument and shall be located such that it is exposed to the same ambient temperature as the scale and load cells. 

The totalizer enclosure shall be constructed of a corrosion-proof non-conductive material.

The totalizer junction box shall hold the pressure transducers, load cell manifold, and associated tubing with provisions for purging each individual load cell hydraulic circuit of air.

The totalizer junction box shall provide sufficient strength to protect the internal components from physical damage.

The totalizer shall be designed for on-site repair. It shall not be necessary to return the totalizer to the factory for any service or repair.

GROUNDING SYSTEM:

The grounding system shall employ a single-point ground in accordance with local regulations. The grounding system shall conform to the manufacturer’s recommendations.

WEIGHT INDICATOR:

The weight indicator shall comply with the appropriate speciﬁcations for a Class IIIL 10,000 division weight indicator as speciﬁed by the National Institute of Standards and Technology Handbook44 and shall have an NTEP Certiﬁcate of Conformance attesting to that compliance.

The weight indicator shall be housed in a stainless steel enclosure and shall include an LCD display visible in all levels of light including total darkness.

The weight indicator shall be manufactured by the scale manufacturer and shall be a Cardinal Scale 225 series Navigator weight indicator or equivalent.

WARRANTY REQUIREMENTS:

The scale manufacturer shall warrant the scale assembly including the deck and components below the deck for a period of ﬁve years; the digital weight indicator, printer and peripheral devices shall be covered for a period of one year. The load cells shall be warranted against lighting and water damage in addition to material and workmanship for the life of the scale.

The manufacturer or its local representative shall present a program of regular maintenance and calibration service. Inspection in said maintenance program shall occur a minimum of once every six months and shall comply with the guidelines set forth by the scale manufacturer, local regulations, and the current edition of the National Institute of Standards and Technology Handbook 44.